Chapter Three

PASSIVE ELECTROMETRIC

CONVERTERS. SOURCES

LOW CURRENT

3-1. HIGH RESISTORS

- Implementation principles

The first high resistance resistor was an ionization chamber operating in the unsaturated mode, where the current is proportional to the applied voltage. Subsequently, other principles for the construction of high -resistance resistors were developed : liquid and solid (bulk and surface conductivity), including the use of black opaque paper, thin metal layers, graphite, ink, soot, lead- silicate glasses [L. 3-2, 3-17].

In electrometry, non-wire high- resistance resistors are used; wire resistors with a resistance of up to 10 12 -10 13 ohms cannot be manufactured. On the basis of a cast microwire , resistors with nominal resistance values up to 10 9 Ohm are made, attempts are being made to manufacture microwire resistors for 10 10 and 10 11 Ohm.

- Composition resistors

The production of high resistance nonwire resistors was mastered after the Second World War. Most of the series resistors are composite lacquer-film resistors. They have a number of advantages. Conductive cross section

element in composite resistors is large, which increases reliability. The production technology does not require complex equipment, expensive materials, high temperature or vacuum processes. During manufacture, it is possible to control the electrical properties of resistors by changing the composition of the composition. In particular, the value and sign of the temperature coefficient can be changed. The disadvantages of composite resistors are the dependence of resistance on voltage, high noise, poor resistance to moisture and temperature, and significant capacitance.

In the USSR, high-resistance resistors of the KVM type are produced - composite evacuated small -sized ones and of the KLM type - composite lacquered small-sized ones. Their properties are regulated by GOST 10686-63 “Constant composite film resistors”, and test methods by GOST 3223-67 “Constant resistors. Test Methods".

In the manufacture of composite resistors, the film composition is applied [L. 3-10] on a ceramic base: porcelain with nominal resistances up to 10 11 Ohm and ultra-porcelain or steatite at large. The composition consists of a phenolic resin and a moderately calcined channel black having a high dispersion and is applied by dipping. The resistance values vary by changing the soot concentration. The final adjustment is made by spiral or longitudinal cutting, and within 10 - 20% - by thermal effects. The resistive element is covered with a protective enamel with a low drying temperature.

KLM resistors are unsealed, and KVM resistors are sealed. In the latter, the resistive element is sealed into a glass canister with a vacuum of 5–10 Pa ( 0.05–0.1 mmHg ). Preliminary training is carried out in a vacuum at 80 o C. The outer surface of the canister is treated with water- repellent agents.

- Basic properties

Characteristics of serially produced resistors with resistance up to 5∙10 14 Ohms are given in Table. 3-1.

Stability. The change in the resistance of composite resistors over time is associated with various 54

Table 3-1

Characteristics of serial high-resistance resistors

| Type of | Manufacturer | Type of resistor | Dimensions, mm |  | Resistance range , Ohm | Permissible deviation, % | Maximum voltage, V, |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Length | Diameter |  |  |  |
| kvm | USSR | Vacuum- sealed | 51 | 5 | 1.5 * 10 7 -10 9 1.2-10 9 —10 12 | 5; ten; twenty; ten; twenty | 100 |
| klm | USSR | non-vacuum | 25-29 | 5.5-7 | 10 7 —10 10 1 .5∙10 10 -10 12 | 5; ten; twenty; ten; twenty | 300 |
| TR142 | Tesla (Czechoslovakia) | Vacuum- sealed | 55 | eight | up to 10 12 | ten | — |
| HWK | WBN People's Enterprise , Teltow (GDR) | non-vacuum | 16-45 | 6-11 | 5-10 6 —10 12 | ten; twenty | 100-3500 |
| RX-1 | Victoreen (USA) | Vacuum- sealed | — | — | 10 8 —10 12 10 12 —10 14 | 2 5; ten | 1000 |
| HSD HSK | Resistance Products Co (USA) | — | 19-75 | — | up to 10 14 | 3; 5; ten | up to 1500 |

| Type of | Manufacturer | type of resistor |
| --- | --- | --- |
| PY3611 PY3623 PY3635 | Welwyn Electrical Lab. Ltd ( England ) | — |
| — | Compagnie Generale de telegraphie _ Jans Fil ( France ) | Vacuum- sealed |
| Stabilme - tal A4 | Radiaco (France) | — |
| SWK 6/70 | Rosental (Germany) | non-vacuum |
| ARC 1.4 | Electronic (Germany) | Vacuumed |

Continuation of the table. 3-1

| Dimensions, mm |  | range , Ohm | Permissible deviation, % | Maximum voltage , V |
| --- | --- | --- | --- | --- |
| length | diameter |  |  |  |
| — | — | 10 7 —10 12 | ten; twenty ten; twenty ten; twenty | 500 1000 2000 |
| 40 | 5 | 10 7 - 5*10 14 | 2; 5; ten | 500 |
| 40 | 6.3 | 10 6 -10 13 | 5; ten; twenty | 1500 |
| — | — | 10 7 -10 11 10 11 —10 13 | ten 25 | 1000 3000 |
| — | — | 5*10 6 - 10 8 10 8 —10 10 10 10 —10 11 | one 5 twenty | 500-8000 |

57

The main conclusions are as follows. Vacuumed resistors have better stability than non-vacuumed ones . Stability tends to increase as the aging time of the resistor increases. The resistance of most (up to 80%) vacuum resistors decreases with aging. Resistors whose resistance decreases over time are more stable than those whose resistance increases. The latter are usually extremely unstable. Among the aged KVM resistors, from 5 to 20% has an instability of less than 0.3 - 0.5% per year for nominal values of 10 9 -10 12 ohms and 0.5 - 1% per year for 10 13 -10 14 ohms. To increase stability, KLM resistors are trained with several

physical phenomena and is complex, depending on the characteristics of the resistor, manufacturing technology, natural and artificial aging and the conditions in which it is located. The main causes of aging are [L. 3-10]: structural changes in the conductive material (crystallization); absorption and evolution of gases, oxidation of the conductive material; irreversible changes in the binding base ( polymerization, volumetric shrinkage, destruction, mechanical destruction); diffusion of the volatile protective coating into the conductive layer, curing of the coating.

Instability is different during storage and under electrical load. The change in the resistance of KVM resistors for 2 years of storage according to GOST 10686-63 does not exceed 5%. Instability of resistors under normal conditions {L. 3-23] the greater the greater the resistance of the resistor (Fig. 3-1). As the temperature rises, the process accelerates, and at 100 o C instability reaches 5% per month. The instability of high-resistance resistors is considered in a number of works [L. 3-28, 3-31]. Main

Rice. 3-1. Dependence of long-term instability on the nominal resistance of the resistor.

temperature cycles from - 40 to +60 o C or kept for several weeks at 50 - 70 o C. KVM resistors are recommended to withstand 100 hours at 100 C , and then 2 - 3 weeks under normal conditions.

The effect of temperature. As the temperature increases, the resistance of the conductive layer decreases, the bonding medium of the conductive layer expands, which increases the resistance, and the conductivity of the insulating coating changes. The resulting dependence of the resistance of composite resistors on temperature can be complex. KVM resistors are designed for operation at temperatures from -60 to +85 o C. For KLM resistors, the lower temperature limit is the same, and the upper one is +100 o C at resistances up to 10 10 Ohm and + 70 o C at large. In table. 3-2 shows the temperature coefficients of resistance (TCR) for resistors KVM, KLM and other types. In practice, the requirements of regulatory documents regarding TMS are not always met. KVM resistors with a resistance of up to 10 11 ohms in most cases meet the requirements of GOST 10686-63, and the TCR of resistors with a resistance of 10 12 ohms is 2-3 times higher than the permissible one.

Influence of voltage. The current density in the contacts between conductor grains of composite resistors can be high. The resulting local overheating causes a change in the resistance of the resistor. As the voltage increases, the resistance of the resistors, in accordance with the sign of the TCR of the conductive layer, usually decreases. With positive TCR , small increases in resistance can also be observed. Nonlinearity can also be caused by voltage dependence of the conduction of gaps or dielectric layers between grains.

The effect of stress is characterized by the stress factor

where R u 1 and R u 2 - resistance at voltages U 1 and U 2

According to GOST 10686-63 U 1 \u003d 10 V and 1 00 V.

In foreign practice, it is customary to use the coefficient

When U 1 = 0, lU 2 we have K ` n \ u003d K n / 0.9 U 2 . The value of K ' n depends on the voltage and characterizes the non-linearity of the resistor only in a given voltage range.

The effect of voltage (Fig. 3-2) is greater, the greater the resistance of the resistor [L. 3-23]. In table. 3-2 shows the voltage coefficients K ` n for resistors of various types. Practical studies of nonlinearity [L. 3-31, 3-35]

Rice. 3-2. Dependence of the voltage coefficient on the nominal resistance of the resistor.

Рис. 3-3. Зависимость коэффициента напряжения от напряжения.

for the most part do not apply to voltages below 1 V, at which resistors in electrometric equipment usually operate. For Victoreen resistors with resistances of 10 10 and 10 11 Ohm dependence K ` h ( u ) has [L. 3-30] is linear (Fig. 3-3), which allows extrapolating it to the region of low stresses (dashed line in Fig. 3-3). Of course, more extensive research in this direction is required.

There are resistors that have normal properties at high voltages and a complete circuit break at low voltages.

Influence of radioactive radiation. The impact of ionizing radiation on high-resistance resistors causes ionization of dielectrics and additional conductivity. Radiation also leads to irreversible measurements of the structure and properties of materials, reducing the x resistance. Composite resistors are more sensitive to the influence of ionizing radiation,

Table 3-2

Temperature coefficient of resistance (TCS) and voltage coefficient K' n high resistance resistors

| resistor type | Executor | tks , %/k | K` n %/AT |
| --- | --- | --- | --- |
| kvm | USSR | From - 0.2 to +0.1 for temperatures from -60 to +20°С; +-0.2 for temperatures from +20 o C up to the maximum | +-0.06 for resistances up to 10 11 Ohm at U = 10÷100 V; from - 0.11 to + 0.06 for resistances more than 1.2 ∙ 10 11 Ohm at U = 10÷100 V |
| klm | USSR | For resistances up to 10 10 ohms, the same as for KVM resistors. For resistances above 10 10 Ohm: from - 0.25 to 0.15 for temperatures from - 60 to + 20 C ; +-0.25 for temperatures from + 20 o C up to maximum | From - 0.17 to +0.06 at U =10÷100 V |
| TR142 | Tesla ( Czechoslovakia) | 0.15 | 0.1 |
| RX-1 | Victoreen (USA) | From - 0.06 to - 0.15 | - 0.02 for resistances 10 10 Ohm; — 0.06 for resistances 10 12 O m; - 0.10 for resistances 10 13 Ohm at U =1÷100 V |
| — | Compagnie Generale de Telegra - pħie Jans Fil ( France ) | From - 0.2 to - 0.4 | — 0.5 at U = 10÷100 V |
| APK1.4 | Electronic (Germany) | 10 9 Ohm: -0.15 10 11 Ohm: - 0.4 | -0.015 -0.04 |

than , for example, metal film or metal oxide . The effect of radiation is reduced by sealing the resistor or by placing it in a plastic shell.

own capacity. Composite resistors have a capacitance of 0.1 - 0.2 pF. It is distributed in nature, and the resistor equivalent circuit can be represented by a series connection of a large number of RC chains. Usually, one RC network is used to describe the behavior of a resistor , but this equivalent circuit is approximate. In order for it to accurately reflect the behavior of the resistor, it is necessary to accept its parameters (7? and C) as frequency dependent.

own interference. Sources of interference in composite resistors have been little studied. Presumably, the constant components of interference are due to thermal emf , contact potential difference, parasitic currents of insulators, etc. The parasitic current of KVM resistors reaches 10 -15 -10 -14 A and depends on the purity of the glass shell. Resistors have thermal and current noise. The source of current (excessive) noise is resistance fluctuations caused by a change in the area of contacts between conductor grains. The level of current noise is proportional to the applied voltage, their frequency spectrum is continuous, and the spectral density depends on frequency approximately according to the law l ∕ f . The level of current noise of a resistor can be characterized by its own noise figure, equal to the ratio of total noise to thermal noise . In GOST 10686-63, the noise of KVM and KLM resistors is not standardized; their experimental determination is a difficult task.

Vibronoise . The resonant frequencies of vibronoise resistors KLM lie in the range of 500 - 550 Hz [L. 3-16 J , their level does not depend on the resistance of the resistor and is 2.5 - 5 mV. The vibration noise of the KVM resistors is much greater, sometimes hundreds or thousands of times. The core of the KVM resistor can, during vibrations, move in the fields created by the charges on nearby insulators , which leads to the appearance of induced charges on the resistive element. If the resistor is energized, then the capacitance of the resistive element is charged a and its change during vibrations also causes the appearance of parasitic signals. To reduce vibration

noise, KVM resistors should be fixed rigidly and shielded from fields. The vibration noise is drastically reduced by centering the resistor core inside the glass bulb with spacers, such as permalloy (Fig. 3-4).

3-2. LOG ELEMENTS

- General information

The current-voltage characteristic of logarithmic elements (LE) for converting small currents into voltage is described by the equation [L. 3-8]

U = a 0 + a 1 log I + ΔU ( I ) , 	(3-1)

Рис. 3-4. Конструкция виброустойчивого резистора.

where a 0 and a 1 - constant coefficients having the dimension of voltage, and under I is understood the value of current , referred to 1 A. The term ΔU ( I ) gives the deviation of the characteristic from the logarithmic one. Electrovacuum and semiconductor LEs satisfy equation (3-1) in the current range up to 6-10 orders of magnitude at the lower limit of the range up to 10-15-10-14 A.

The current-voltage characteristic of a particular LE can be represented by an infinite set of expressions of the form (3-1) with different values of the parameters a 0 , a 1 and ΔU ( I ). Such an expression from this set, at which the maximum value of the term ΔU ( I ) is minimal , is called the logarithmic characteristic of the LE. Current range for which a 0 , a 1 are set and ΔU ( I ), is called the range of the logarithmic characteristic . The coefficient α 0 , equal to the voltage on the LE at a current of 1 A, is called the displacement of the logarithmic characteristic, its change shifts the characteristic in the direction of the voltage axis. Coefficient a 1 is equal to the voltage increment on the LE when the current changes per decade and is called the steepness of the logarithmic characteristic. The term ΔU ( I ) represents the voltage of the non-linearity characteristic.

The dynamic resistance of the LE is found by differentiating the equation (3-1), it is inversely proportional to input current

The static resistance of the LE also depends on the current through the element, it is maximum at

The main source of error in logarithmic current-to-voltage converters is the non-linearity ∆ U ( I ). From equation (3-2) it follows that the relative error of current conversion with the help of LE is associated with its nonlinearity by the expression

i.e., it is constant when the absolute error of the voltage conversion is constant. To reduce the relative error, it is necessary to minimize the ratio ∆ U / a 1 .

A significant source of error is the effect of destabilizing factors, of which the main ones are temperature, and for electrovacuum LEs, also the parameters of the electrical regime. Wherein all parameters of the LE characteristic can be changed. The time drift of the logarithmic characteristic also manifests itself in the change of these parameters.

- Electrovacuum logarithmic elements

Parameters of the logarithmic characteristic. In electrovacuum LEs , the current-voltage characteristic is determined by the distribution (according to Maxwell) of the velocities of electrons emitted thermionic cathode at negative and small positive voltages on the anode, and has the form:

where U is the potential difference between the anode and cathode of the diode; I is the current through the diode; I 0 - current through the diode at U = 0; K- _ Boltzmann's constant; T is the absolute temperature of the cathode; q is the charge of an electron.

63