© Издательство «Энергия», 1976 г.

| 30306-054 |  |
| --- | --- |
| И 051(01)-76 | - 258-75 |

6П2.1.081

6 Π 2.1.081 AND 48

UDC 621.317.3.01

Ilyukovich A. M.

And 48 Electrometry technique. M., "Energy", 1976. 400 p. from ill.

The book discusses methods and means for measuring small electrical signals in high-resistance circuits and the resistance of such circuits at low levels of constant or slowly varying currents (up to 10 -16 -10 -17 A). General information about electrometric measurement objects, about electrometric equipment and its elements is given. The principles of construction and properties of various types of electrometric voltage amplifiers are described in detail . Issues related to the construction of electrometric meters for voltages, charges , currents, resistances, parameters of capacitive objects at infralow frequencies are analyzed. The methods and means of verification of electrometric equipment are described, the features of its operation are touched upon.

The book is intended for specialists in electrical measuring technology involved in the creation and production of electrometric equipment, as well as for people who use it in practice.

FOREWORD

The development of modern science and technology is characterized by an increasing differentiation of individual areas . This trend also extends to electrical measuring technology, within which, relatively recently, another new direction has emerged - the so-called electrometry.

The appearance of the very concept of "electrometer" dates back to the middle of the 18th century. Its origin is associated with the work of Acad. G. V. Richman, who created the world's first device that had all the main elements of a modern electrical measuring instrument. In this device, as well as in subsequent systems of a similar purpose , the forces of electrostatic attraction were used to evaluate the studied electrical phenomena. At that time, it was not yet clear that electrical phenomena are diverse and characterized by a large number of interrelated parameters, so the introduction of the concept of "electrometer" was quite natural.

In the 19th century, when other types of electrical measuring instruments were created (galvanometer, ammeter, voltmeter, electricity meter, etc.), the concept of “electrometer” was somewhat refined and they began to designate highly sensitive devices of an electrostatic system with a very large input resistance , i.e. with very low input current.

By the end of the XIX century. electrostatic electrometers reached a high degree of perfection and were used as one of the main means of electrical measurements in experimental physics. Maxwell, and then other authors, developed their detailed theory. Using electrometers, it turned out to be possible to measure not only low voltages in high-resistance circuits, but also small electric currents and charges of ionization chambers and other objects, as well as high resistances. Experimental studies carried out with the use of electrometers by Rutherford, the Curies and other leading scientists served as the basis for fundamental discoveries in nuclear physics.

In the 1930s, along with electrostatic electrometers, vacuum tube amplifiers with a low grid current and high input resistance, which were called tube electrometers, began to be used . Immediately after the Second World War, electrometers based on the so-called dynamic capacitor appeared, and then electrometers based on field-effect transistors with a MOS structure and other elements. In the last 20-30 years, on the basis of various electrometers, a large number of various instruments with wide metrological, functional and operational capabilities have been created; these devices are usually combined with the concept of "electrometric equipment".

During this time, the areas of application of electrometric equipment for scientific research have expanded many times and it has been introduced into industry , energy, health care, etc. The level of development of modern electrometric equipment largely determines the progress of a number of important areas of science and technology: the technique of physical experiment, technology of dielectrics and high-resistance semiconductors, microelectronics, ionizing radiation technology, mass spectrometry, vacuummetry, gas chromatography, low-frequency piezometry, electrochemistry, pH-metry, biology (intracellular research), cosmic radiation research and many others.

research, development, production, and use of electrometric equipment. The number of types of serial electrometric equipment in our country is in the tens, and if we take into account the measuring instruments in which such equipment is used in the form of built-in units, then hundreds. Every year a large number of works on electrometric methods and measuring instruments are published in the world .

Thus, the electrostatic electrometer, which has served well in the technique of physical experimentation, has been replaced by a new variety of apparatus, with the help of which not so much laboratory as industrial problems are solved. As part of electrical measuring technology stood out another direction. It is usually called electrometry ( there is also a textual interpretation of the term “electrometry”, when it refers to all electrical measuring equipment in general; we will use this term only in its conventional meaning, the origin of which was explained above). Modern electrometry is characterized by independent areas of application and its own electrometric ranges of measured parameters. The physical processes that determine the metrological characteristics of electrometric measuring instruments are in many respects specific and not typical for other areas of electrical measuring technology. Accordingly, the principles for constructing such measuring instruments and their elements are also specific . Methods and means for metrological assurance of the unity and correctness of measurements in the electrometric range are distinguished by great originality . The rules for operating electrometric equipment are also different from those generally accepted for other electrical measuring instruments .

At the same time, the experience accumulated in the field of electrometry has not yet been generalized, which hinders its further development. Many interesting works have been published on certain types of electrometric equipment, including a number of monographs. However, there is no general consideration of all methods and means of electrometric measurements in their natural interconnection from a unified scientific and technical position and with the same depth of study of individual issues. The solution of this problem is the subject of this work.

What has been said above about the originality of electrometry as a direction of electrical measuring technology, of course, does not mean that electrometry develops outside the scope of electrical measuring technology as a whole. Many measurement methods and principles for constructing measuring instruments are successfully borrowed in whole or in part. It was the development of electrical measuring technology that created the necessary basis for the modern development of electrometry. In this work, however, we will be primarily interested in precisely those questions that are specific to the methods and means of electrometry. It is assumed that the reader is familiar with the basics of electrical measurement technology, especially with direct current measurements.

When writing the book, the author equally relied on numerous publications on individual issues of electrometry and on more than a decade of experience of the VNIIFTRI electrometry laboratory he leads. V. M. Borzov, L. A. Vsevolozhsky, I. V. Dzyubenko, L. V. Klemina, V. A. Kulikov, I. P. Mikhailova, T. O. Petrov, A. E. Strelichev, Yu. A. Tsygankov and others, the author asks to accept his deep gratitude, since without their creative activity the appearance of this work would have been impossible. The author is grateful to a large detachment of colleagues from other organizations for their constant friendly assistance in working on the problems of electrometry and the manuscript of this book. Made by the reviewer Dr. Techn., Sciences D. E. Polonnikov's remarks made it possible to significantly improve the manuscript of the book.

As the first attempt in the field of a generalized consideration of electrometry, the book is not without shortcomings. All comments that readers will take the trouble to send to the address: Moscow, 113114, Shlyuzovaya nab., 10, Energiya publishing house, will be gratefully and carefully studied.

Author

Chapter first

GENERAL INFORMATION

ABOUT ELECTROMETRIC METHODS

AND INSTRUMENTS

AND ELECTROMETRIC MEASUREMENT OBJECTS

1-1. GENERAL INFORMATION ABOUT ELECTROMETRIC METHODS AND MEASUREMENT INSTRUMENTS

- Distinguishing properties

The most characteristic feature of electrometric equipment, which distinguishes it from other types of electrical measuring equipment, is the specific nature of the sources of intrinsic noise, which determine the realistically achievable sensitivity threshold. These include parasitic currents of insulators, parasitic currents from background radiation and radioactive contamination of materials, electrostatic fields created by charges on insulators and contact potential difference, etc.

All these sources of self-interference appear only at low currents. Consequently, the most common feature of all types of electrometric measuring instruments is a low current level in the input measuring circuits (of the order of 10 -17 -10 -10 A). This feature determines many features of electrometric equipment.

First, small currents exceed the thermal noise currents and, therefore, can be detected against the noise background only in very high-resistance circuits. Root mean square value of intrinsic current of thermal noise of an object with resistance R in the frequency band Δf

равно:

where K = 1.38∙10 -23 J/K is the Boltzmann constant; T is the absolute temperature (see § 1-3).

If, for example, at ∆ f = 1 Hz it is required to obtain at room temperature (T ≈300 K) for some circuit √ i 2 t ≤10 -15 A , to it is necessary that the resistance of this circuit R ≥1.5∙ 10 10 ohms. Thus, the input circuits of electrometric equipment are high-resistance.

Second, small currents can exceed thermal noise currents only in a not very wide frequency range . The thermal noise current increases proportionally to √ Δf and for large ∆ f can be significant even at high circuit resistances. In addition, if the resistance R circuit is large, then even small shunting R capacitances lead to a significant bandwidth limitation of such a circuit. Thus, the characteristic features of electrometric equipment manifest themselves to a greater extent, the smaller its bandwidth.

Note that electrometric measuring instruments operate only with constant and slowly varying signals. It is in this frequency range that most of the indicated sources of intrinsic interference operate and the features of electrometric equipment appear.

All of the above applies to both electrometric measuring instruments and measurement objects of the electrometric range, the latter are also high-resistance and operate with constant or slowly changing signals. The sources of intrinsic interference of such objects are determined by the low level of current in them and, as a consequence, their high resistance.

- Main types

Electrometric measuring instruments, like any other [L. 1-25], according to their purpose, they are divided into measures, measuring transducers, measuring instruments and complex measuring devices. The latter have not yet received application in electrometry, so further consideration will be limited to the first three groups. We also don't boo- 8

), i.e., the sum of variables, slowly changing and constant components. We will not take into account the input capacitance of the meter, since the transient or frequency properties of the meter are not of interest to us yet .

An equivalent circuit for measuring voltages, compiled taking into account what has been said, is shown in fig. 1-1, where E is an ideal (with an infinitely large input resistance and no intrinsic noise) electrometric voltage meter.

Indications of the electrometric meter E

must make a clear distinction between measuring transducers and instruments, since in electrometry most often the difference between them consists only in the presence of built-in output pointers in instruments. First of all, devices and transducers will be considered, and the measures will be considered separately in the sections on the verification of meters of different quantities.

On the basis of measured or reproducible quantities, electrometric measuring instruments are divided [L. 1-13 - 1-15] on measuring instruments for voltages, currents, charges (signal parameters), resistances and capacitances (circuit parameters).

c) Construction principles

Let us consider the main features of measurements of each of the parameters of signals and parameters of circuits in the electrometric range. Let's start with measuring voltages. The signal source will be characterized by the internal resistance R i and useful signal U x , and the meter - input resistance R in and the level of self-interference in voltage e p and current i p

(see § 4-2). By self- interference we mean the totality of noise, drift, and zero offsets (in terms of voltage and current, respectively

9

An indispensable condition for the correctness of measurements is R in>> R i . At R i = 10 12 ÷ 10 13 Ohm, which can

Rice. 1-1. Equivalent circuit for measuring voltages.

иметь место на практике, необходимо, чтобы Rbx≥1014÷1016Ом. Эти значения приближаются к предельно возможным, т. е. к тем, которые обеспечиваются при использовании в качестве изоляции воздушных промежутков. При соблюдении условия Rnχ^Ri показания электрометрического измерителя

i.e., the relative measurement error caused by the meter’s own noise is ( e p + i p Ri )∕ U x and for

its decrease, it is necessary to reduce e p and i p . Thus, an electrometric voltage meter should have as much input impedance as possible and possibly less intrinsic voltage and current noise.

To explain the principle of operation of electrometric current meters, it is advisable to use the concepts of current-controlled meters and voltage-controlled meters . In the first case, the meter responds to the presence and level of current at the input. A typical example is the devices of the magnetoelectric system. In the second case, the meter reacts to the presence and level of voltage at the input. An example would be the devices of the electrostatic system. In meters controlled by current, the input circuit is most often represented by a winding through which the measured current is passed. For design and technological reasons, the winding resistance does not exceed hundreds or thousands of ohms, and the level of thermal current noise is high. Such devices are not suitable for measuring low currents. In voltage-controlled meters, the input resistance can be very large (up to infinity), and the thermal current noise is very small.

In principle, current-controlled meters can also be sufficiently high-resistance (up to 10 9 -10 10 ohms) to be used in the electrometric range. For example, in amplifiers with an input stage based on bipolar transistors, the control signal is the current, however, it is possible to obtain a large input resistance and provide a high current sensitivity threshold - up to 8∙ 10 A. In view of the high resistance of the input circuits, all the specific features of the electrometric equipment are preserved in this case as well. ten

However, the maximum possible current sensitivity threshold can only be achieved with the help of voltage-controlled meters, so in the future only they will be considered.

At the same time, to measure a small direct current, it must be converted into voltage, and then this voltage must be determined with an electrometric meter. There are two types of small current-to-voltage converters: a resistor with a large (10 8 -10 9 ohm and higher) resistance R and a C capacitor with low leakage. The voltage across the resistor is proportional to the measured current I x , namely

capacitor voltage

Rice. 1-2. Equivalent circuit for measuring currents.

where Δt is the charge time, i.e., the capacitor converts the current into the rate of voltage change ( I x = CΔU ∕ Δt ) . Instead of a resistor, a high-resistance resistor element with a logarithmic relationship between voltage and current is sometimes used, which is often useful. There are no other ways to convert small currents into voltage (and cannot exist , since other elements

there are no cops of electrical circuits with a resistance large enough for the purposes of electrometry, except for a resistor and a capacitor).

Consider the requirements for an electrometric voltage meter for measuring low currents by the resistor method. The equivalent circuit for measuring currents is shown in fig. 1-2, where R 0 is the resistance of the resistor -measure. The voltage measured by the electrometer is:

where R ς = R in R 0 ∕( R in + R 0 ). The condition for correct measurement is R at x >> R 0 . Using

an electrometric amplifier (EMA) instead of a simple electrometer, the equivalent input impedance of the EMA can be made very small due to the deep parallel negative feedback (NFE) , but this does not contradict the last statement. For an EMU, the very high resistance requirement will apply to the area between the input and output potential terminals. If the condition is met, the error is equal to:

To reduce it, it is necessary to increase the resistance of the resistor-measure (respectively, R in ) and reduce the intrinsic interference in voltage and current, i.e., the requirements for an electrometric voltage meter are the same as in the previous case. It is easy to show that similar requirements are also obtained when measuring small currents using a capacitor.

When measuring small , slowly changing charges, it is also necessary that the current level in the input circuits be small. Otherwise, the rate of change of the measured charge due to the presence of this current ΔQ x ∕ dt = I ( t ) will be significant and one can

control only fast processes.

Low-resistance meters, for example, a ballistic galvanometer , do not allow measuring small charges, since the thermal noise in the current, and , consequently, in the charge, is large. The most sensitive ballistic galvanometer type M17 / 13 has a constant of only 0.5 * 1 0 -9 C-m / mm. To measure small charges, they must be converted into voltage, and then this voltage should be determined with an electrometric meter . For these purposes, capacitors with high leakage resistance are used. Voltage Uc _ on a capacitor with a capacity of C 0 is related to the measured charge Q x known expression U c = Q x / C 0 .

When measuring charges, parallel to the capacitance of the source C i , the charge of which is measured, is connected

12

Rice. 1-3. Equivalent circuit for measuring charges.

measuring capacitance C 0 (Fig. 1-3). Voltage at the electrometric meter where t is the measurement time.

The condition for the correctness of measurements is R in * ( C i + C 0 ) >> t . In this case, the charge leakage over the -resistance Rin during the measurement time can be neglected, and then it is not necessary to know the capacitances of the source and the connecting line. Under these conditions, the measurement error is equal to ( i p t + e p C 0 )∕ Q x and to reduce it, it is necessary to reduce the interference e p and i p , i.e., the same requirements are imposed on the electrometric voltage meter as in the previous cases.

When measuring the parameters of circuits, the object is not a source of energy, and it is necessary to use an auxiliary source, measuring the parameters of the signals that arise in this case in the controlled circuit. When measuring resistance, it is necessary to obtain information about the voltage on the object and the current through it. The voltage meter must be electrometric, since the resistance of the object can be high. The currents through the object are small, and the current meter must also be electrometric. In special cases, one of these meters can be dispensed with, for example, if you set the voltage on an object and measure the current through it, or you set the current through the object and measure the voltage across it. A similar situation occurs when measuring small capacitances at infra-low frequencies.

d) Basic elements

So, for any measurements in the electrometric range, an electrometric voltage meter with a high input resistance and low intrinsic voltage and current noise is used.

13

To measure small currents and charges, in addition, a passive electrometric converter of the measured value into voltage is required. If measurements of high resistances and small capacitances are reduced to measuring currents or charges, then such a converter is also necessary in this case. When measuring charges , only a capacitor can serve as a converter (meaning the case when the charge is stored during the measurement; if this is not required, then the charge can be passed through the resistor and the charge can be found from the integral value of the dependence of the current through the resistor on time), and when measuring currents are also resistive elements : high-resistance resistors or logarithmic elements. All of them have received practical application. Passive electrometric transducers belong to the category of passive two-pole [L. 1-12].

The error of the main characteristics of such transducers (capacitance for a capacitor, resistance for a resistor, and current-voltage dependence for a logarithmic element) is fully included in the measurement error.

Therefore, this characteristic must be precisely defined, stable over time and under changing environmental conditions. Capacitors have the best properties in this regard, and logarithmic elements have the worst.

The capacitance of the resistive elements and the leakage resistance of the capacitors are residual (not the main ) parameters. They define the transient characteristics of the converters: the time constant for resistive elements and the maximum allowed conversion time for capacitors.

Passive converters may have their own interference. Like all passive two-terminal networks, it is enough to characterize them with one source of their own interference: either by voltage or by current [L. 1-12]. The noise of passive electrometric transducers has a direct effect on the measurement error, and it is desirable that they be small.

Some other special elements are also used in electrometric equipment. First of all, these are sources of low currents. They can be used, as mentioned above, when measuring high resistances, as well as in some other cases, to

the errors of their main characteristic - the output current - are subject to stringent requirements.

In addition, structural elements of a special design and execution are used in electrometric equipment : insulators, switches, relays, etc. The general requirements for them are high leakage resistance and a low level of self-interference. e) Limits of the electrometric range

The sensitivity threshold of electrometric meters is determined by the level of their own additive noise. Currently, the best electrometric voltage meters have a sensitivity threshold of approximately 2 μV, current meters 10 -17 A, charge meters 10 - l 6 C, resistance meters 10 18 -10 19 Ohm. The theoretical threshold of sensitivity lies even lower (§ 1-3).

As the currents at the output of controlled electrometric objects increase, the resistances of the input circuits of the electrometric equipment can be chosen to be more and more low-resistance, and the effect of intrinsic additive noise specific to the electrometric range decreases. Tasks of a different nature come to the fore.

The boundary separating electrometric equipment from non-electrometric, of course, is not clear . It depends on the level of currents in the input circuits, the required measurement accuracy, the frequency range, etc. Usually, even at currents of 10 -11 -10 -10 A, the specific features of electrometric equipment do not appear. In practice, electrometric meters often provide measurement ranges at which currents are significantly higher than the specified level and sometimes amount to 10 -4 -10 -2 A. Thus, a very wide operating range is realized. This is convenient , however, in essence, the use of electrometric measurement methods becomes justified only at currents below 10 -10 - 10 -8 And when current-controlled meters (for example, photogalvanometric amplifiers) cease to be useful. We will consider these values as the upper limit of the electrometric range when measuring currents.

The upper limit of the electrometric range when measuring voltages is also determined by the level of currents

in the input circuits of the meters. In this case, the resistance of the signal source should not be less than 10 5 -10 6 ohms, otherwise the thermal noise current will be too large. Accordingly, the input resistance of the meter should not be lower than 10 7 - 10 8 Ohm. In each specific case, it is easy to determine whether this or that problem is electrometric or not. For example, to measure a voltage of 1 mV from a source with an internal resistance of 10 8 ohms , a meter with an input resistance of the order of 10 10 ohms is required, i.e., the current will be 10 -13 A, and the use of an electrometric meter is certainly necessary. If in the same source the voltage is 100 V, then the current will be 10-8 And you can do without an electrometric meter.

When determining the boundaries of the electrometric range of charge meters, it is necessary to take into account not only the level of measured charges, but also the time constant T of the controlled circuit. The charge in a circuit with time constant T changes exponentially, the initial segment of which is described by the equation

where Q0 _ is the charge at Δt =0.

Therefore, charge leakage ΔQ = Q 0 Δt ∕ T n leakage current I ( t ) =∆ Q ∕ Δt = Q 0 ∕ t For example, to measure the charge yes Q 0 = l 0 -8 Cl at T \u003d 10 6 s requires electrometric equipment ( I =10 -14 A), and at Q 0 \u003d 10 -10 Cl and T \ u003d \u003d 0.01 s , you can do without it ( I \u003d 10 -8 BUT).

Similarly, the current level can be used to determine the need for electrometric equipment in the case of measuring resistances or capacitances.

1-2. GENERAL INFORMATION ABOUT ELECTROMETRIC MEASUREMENT OBJECTS

a) Preliminary remarks

The objects that require electrometric equipment to measure their parameters are very diverse. On the basis of the measured value, they can be divided into sources of voltages, currents and charges, resistive objects and capacitive objects. Below we will consider the main areas of application of each type of electrometric equipment and the primary measuring transducers used in this case. This will make it possible to present 16

to develop the requirements for various types of electrometric equipment in relation to metrological characteristics, functional features, operating conditions, etc.

- Voltage sources

Sources of low constant and slowly varying voltages with high internal resistance can be subdivided into electrophysical, electrochemical, and electrobiological sources based on the processes that cause the appearance of a signal.

Various electrophysical sources of voltages in the electrometric range have to be encountered in physical experiments. Examples are measurements of galvanomagnetic, thermoelectric, piezoelectric, and other effects in high-resistance materials, in particular, at very low temperatures. To solve the most typical problems , specialized equipment is produced, for example, cameras (with highly sensitive electrometric voltage meters) for measuring the Hall effect in high-resistance semiconductors. When studying this kind of electrophysical objects, meters with direct reading or self-recording meters with a sensitivity threshold at the level of units - tens of microvolts, sometimes with high accuracy, are usually required; high-speed equipment is usually not needed.

A widespread electrophysical object is

electrostatic field. Electrometric instruments are used to measure electrostatic potentials in meteorological studies, the ranges of measured voltages are 1-100 V , and the operating conditions are extremely difficult . The need to measure the potentials of electrostatic fields often arises in various industries and transport in the control of explosion and fire safety (oil transport,

film industry, etc.), as well as in some technological processes in paper production, the textile industry, etc. High sensitivity is necessary when measuring fields in the process of studying the electrification of materials, when studying electrophotographic processes, and, especially, when measuring the parameters of fields created by biological objects . Often, when measuring fields, it is required that the meter has a small input capacitance; otherwise, the charge induced at the input of the meter will strongly distort the field under study . The input resistance of potential meters, as a rule, should reach 10 15 -1 0 l 6 Ohm, to minimize the leakage of induced charge through the input circuits to the minimum possible. In some cases, when measuring fields, high-speed meters are used.

Of the high resistance electrochemical objects that are sources of low constant voltages, glass electrodes are the most common for measuring the concentration of aqueous solutions using the electrochemical potential at the boundary of electrodes immersed in the test solution. These instruments are called pH meters.

necessary in the chemical, metallurgical, textile, pulp and paper, pharmaceutical and food industries , in medicine, biology, soil science, etc. Industrial

The production of various pH meters has been established in all developed countries.

The internal resistance of a glass electrode depends on the design and materials, usually it ranges from units to hundreds of megohms. Therefore, the meter of constant electrical potentials on the electrode must have an input resistance of up to 10 11 -10 12 ohms. The range of potentials encountered in pH metering ranges from a few millivolts to a few volts for industrial pH meters, while lower levels are used in laboratory and special pH meters. If the range of measured potentials is wide, and switching the measurement limits of instruments is undesirable or impossible, electrometric voltage meters with a logarithmic characteristic ( U out \ u003d a + b * lgU x ) may be useful. The processes involved in pH metering are usually slow, and potential meters in pH meters are not subject to strict speed requirements. Industrial pH meters for process control are most often self-recording. The measurement error does not exceed 0.5%. In recent years, digital reading pH meters have begun to appear.

In electrobiological studies, electrometric voltage meters are used, in particular, to measure the partial pressure of carbon dioxide in the blood and to study bioelectrical phenomena in plant and animal tissues. The problem of recording the electrical activity of biological preparations using glass ultramicroelectrodes is especially common . For these purposes, serial production of devices has begun abroad . In such intracellular studies, the potentials generated by the cell range from about 1 mV (initial potential) to 200 mV (action potential). Due to its small diameter and the presence of an electrochemical transition, the resistance of a glass electrode is high (tens to hundreds of megohms), and the output capacitance of microelectrodes lies in the range from units to tens of picofarads. An amplifier for working with microelectrodes must have an input impedance of up to 10-10 ohms and a sensitivity threshold of fractions of a millivolt. Since the recorded potentials can change at a high speed, the amplifier must have a high speed. In addition, it must have a small parasitic current at the input - less than 10 -10 A, otherwise there will be additional voltage drops that distort the measurement results , and the cell under study will be irritated by this current and may change its characteristics.

The above consideration shows that, depending on the type of object, electrometric meters have to measure voltages from a few microvolts to tens or hundreds of volts. The required input impedance can be up to 10 15 - 10 16 ohms. A wide range of requirements can be placed on frequency response, accuracy and functionality , and operating conditions can be not only laboratory, but also industrial and field.

- Current sources

The vast majority of electrometric objects of measurement , the output signal of which is current, belongs to the category of electrophysical objects. Among them, one can distinguish gas (including air), vacuum, solid-state and liquid objects.

The most common gas (air) source of low currents is the ionization chamber. The internal resistance of the chamber in saturation mode is very high, and it can be considered as a current generator close to ideal. The current at the output of the chamber is proportional to the power of the radioactive radiation acting on it. Ionization chambers are used to measure ionizing radiation and some other non- electric quantities. The areas of application of ionization chambers are extremely diverse: physical, biophysical and biochemical research, nuclear power engineering, metrological work in the field of ionizing radiation, geological exploration and exploitation of uranium and other deposits, radiobeam therapy, collective and individual dosimetric monitoring, including in military field conditions and on spacecraft, space research of radioactive emissions, measurement of non-electric quantities (non-contact thickness gauges, density meters, etc.), etc.

Depending on the power of controlled radiation, the volume of the chamber, etc., the chamber current can range from 10 -17 to 10 -7 A. When cameras are used for metrological purposes, the current measurement error should not exceed tenths or even hundredths of a percent; in other cases, such stringent accuracy requirements are usually not imposed. High requirements for the fast action of current meters are also usually not put forward, and it is often permissible to measure the average value of the current over a certain period of time. In some cases, for example, when monitoring and automatically controlling nuclear reactors or when registering ionizing radiation in outer space, the current meter must be fast.

The requirements for the recording devices of chamber current meters can be diverse. When monitoring long processes, self-recording electrometric current meters are used. For installation on space vehicles, as well as in some other cases, meters with an output signal in the form of a number of pulses or a frequency turn out to be very convenient. Increasingly, small current meters with a digital readout and with access to a digital computer are being used. Often used devices with a logarithmic characteristic, providing a wide operating range without switching measurement limits. They are used to determine the power and period of nuclear reactors, control the content of tritium in the atmosphere, measure the power of radiation in outer space and ground conditions, etc.

When working with ionization chambers, it is often necessary to measure the ratio or difference of two small currents. Ratio meters, for example, are used in the verification of x-ray dosimeters, and difference meters are used in continuous radioactive thickness gauges.

The use of an ionization chamber is useful not only for measuring the parameters of radioactive radiation. The current at the output of the chamber is proportional to the pressure of the gas in p, and this makes it possible to manufacture vacuum gauges with very high characteristics.

— ionization radioactive vacuum gauges allow measuring pressure up to 10 -6 — 10 -8 Pa (10 -8 - 10 -10 mmHg Art.). At the same time, the requirements for high sensitivity and a wide operating range are imposed 	on the current meter at the output of the converter (sometimes meters with a logarithmic scale are used); high accuracy and speed are not required. Radioactive vacuum gauges are widely produced commercially.

Ionization chambers operating in a non-saturated mode, when the output current is proportional to the applied voltage, make it possible to measure the intensity of electrostatic fields. In this case, the ionizing effect is constant and the current is proportional to the field strength. For air ionization, it is possible to use not only built-in radioactive sources, but also high temperature (flame ionizers ), corona discharge, etc. Abroad, tension meters with a radioactive ionizer are commercially available. In this case, the current meter does not require high sensitivity , accuracy, and speed.

Ionization converters are widely used in gas chromatographs (gas analyzers). In them, the studied gas mixture is ionized using a radioactive or flame ionizer. The current at the output of the detector is proportional to the amount of the investigated component. The components of the analyzed gas are preliminarily separated and fed into the detector in a predetermined sequence. Registration of currents at the output of the detector using a recorder gives a chromatogram; the areas of the peaks on the chromatograms are proportional to the amount of each of the components of the analyzed gas mixture.

Gas chromatographs with electric current meters are widely produced commercially. In industrial chromatographs, and even more so in equipment installed on space vehicles , the operating conditions are difficult. Current meters in chromatographs do not require particularly high sensitivity and speed, but the error should be small. It is sometimes desirable to have a wide operating range without switching limits (i.e. logarithmic response) and direct reading of the cumulative values of the chromatogram peaks . The latter is implemented with the help of electrometric meters of charges received at the collector of the ionization detector.

In all cases considered, the gas under study is forcibly ionized. In electrophysical studies of the atmosphere, in medicine, and in a number of other cases, there is a need to determine the degree of ionization of the air around us. This problem is solved with the help of aspiration counters, which are used to determine the concentration and mobility of air ions. In aspiration counters, the air being examined is forcibly blown between the electrodes, to which an auxiliary constant voltage is applied. In this case, part a of the eroions is deflected to such an extent that it falls on the collecting electrode. The resulting current is proportional to the concentration of air ions. Finding the concentration at different auxiliary voltages and air blowing speeds, one can determine the spectral characteristics of air ions, i.e., the dependence of concentration on mobility (speed of movement in a unit electric field); the latter is determined by the size of the air ion.

In air ion counters, electrometric meters are used for both instantaneous and average currents over a certain period of time. Some modifications of counters use

differential meters that determine the difference between the main current and the interference current. The measured currents usually lie in the range of 10 -14 -10 -12 A, sometimes it is required to measure currents up to 10 -16 -10 -15 A. In some cases, a continuous recording of the results is made. Meter accuracy requirements can be high; usually the speed of the device is not significant. Operating conditions can be both laboratory and field. Aspirating ion counters are commercially available.

A peculiar gas source of low currents is the atmosphere. In the practice of meteorological observations and research , the problems of measuring electric currents in the atmosphere, including the vertical conduction current, total vertical current, and precipitation current, are standard and very widespread. In this case , it is necessary to control, and more often to register , stationary and slowly changing processes; very high sensitivity and a wide operating range are often required; there are no high requirements for measurement accuracy; however, operating conditions tend to be severe.

Vacuum systems, the output of which is a small constant or slowly varying current, are as common as the gas systems just discussed. In vacuum systems, a beam of charged particles, the intensity of which must be determined, is collected on a collector or probe, and the current in the circuit is determined using an electrometric meter. In some cases, these currents can be very small. Since the energy of the particles in the beam is usually high, the current can be preliminarily amplified with the help of an electron multiplier, which is often done. -

Mass spectrometers are one of the most common low current vacuum sources. They serve to analyze the chemical composition of gas mixtures by obtaining the mass spectrum of all positively charged ions in the mixture and make it possible to detect not only individual elements, but also their isotopes. The analyte is converted into a gaseous form and ionized in one way or another; a beam is formed from the resulting ions using an ion optical system and an accelerating electric field , which is then decomposed into components that differ in the ratio of the ion mass to their charge.

If the collector currents are measured directly (without an electron multiplier) , then very high sensitivity requirements are imposed on the electrometric meter. High accuracy and speed are also usually required. It is sometimes desirable to have a wide operating range without limit switching (logarithmic ). There are cases when only the variable low- frequency component of the ion current is determined, sometimes it is necessary to know the ratio of two ion currents. The operating conditions of mass spectrometers are most often laboratory. Mass spectrometers with highly sensitive low current meters are commercially available.

Similar requirements for electrometric current meters are also imposed in helium mass spectrometric leak detectors. In them

a highly sensitive magnetic mass spectrometer is tuned to detect helium passing through the vacuum connections of the test object. Unlike conventional mass spectrometers, in this case there is no need for high accuracy and fast

action. Mass spectrometric leak detectors are also commercially available.

In cases where the currents in the beams of mass spectrometers, cesium frequency references, etc. are preliminarily amplified by electron multipliers, the requirements for the sensitivity of low current meters are significantly reduced. The dark current of electron multipliers is 10 -13 -10 -11 A. Measurements of small currents of positively or negatively charged particles in vacuum are also carried out during probe studies of beams in cyclotrons, betatrons and other charged particle accelerators, as well as in electron microscopes, etc. Usually there are no high requirements for the sensitivity of electrometric meters. The exception is probe measurements of the intensity of scattered electrons in electron diffraction, when, for the purpose of structural analysis of substances, the diffraction pattern that arises during the interaction of penetrating radiation with the substance under study is determined. In this case, electrometric meters with extremely high sensitivity are required.

A widely used low current output vacuum device is the photocell. Photocells are used in physical research, in astronomy, in meters of light characteristics, lighting machines, etc. The variety of areas of their use predetermines the variety of requirements for low current meters used with them in terms of sensitivity, accuracy, speed, functionality, operating conditions, etc. Extremely high sensitivity in this case is usually not needed, since the dark current of vacuum photocells, as a rule, lies at a level not lower than 10 -14 -10 -12 A. Approximately the same requirements are imposed on low current meters, used with photomultiplier tubes, which have approximately the same characteristics and applications as photocells.

A relatively new and still little used vacuum current source is a vacuum chamber for measuring high dose rates of X-ray and γ - radiation. In it, the electrodes are separated by a vacuum gap, and when they are irradiated, a differential current is formed, which is measured by an electrometric device.

A peculiar vacuum source of small currents is space, in which there is always a stream of fast, predominantly charged particles (solar wind). To measure the intensity and spectral characteristics of the solar wind, ion traps, spectrum analyzers, etc. are used. The currents from the collectors of these devices are small, and electrometric meters are needed to determine them. The specific requirements for low current meters in this case are harsh operating conditions and wide range. Of no small importance is the method of presenting information, the most suitable are meters with the conversion of small currents into frequency.

The third category of objects, when determining the properties of which it is necessary to determine low currents, are solid objects: dielectrics and high-resistance semiconductors. They are encountered in practice as often as gas and vacuum ones, however, they are usually quite linear systems and therefore are not considered as current sources, but are characterized by 22

resistance. The exceptions are some physical objects and a number of elements of electronic technology. In particular, electrometric instruments are used to measure the reverse currents of semiconductor diodes and triodes, gate currents of field-effect transistors, input currents of integrated amplifiers, grid currents of vacuum tubes, etc. Sometimes specialized devices are produced for these purposes. The requirements for such devices in terms of sensitivity can be varied, most often the accuracy and speed should not be high, but the operating conditions. laboratory or industrial.

Liquid objects adequate to a source of low currents are much rarer than gas, vacuum, and solid objects. An example is a polarographic cell for the electrochemical determination of the partial pressure in an electrolyte. There are no strict requirements for the sensitivity of the electrometric current meter and its other characteristics.

The review of the most common cases of measuring low currents shows the diversity of requirements for such devices at the present stage of development. In many cases, instruments of extremely high sensitivity are required. The measurement error can range from tenths of a percent during metrological work to tens of percent during evaluative measurements; high speed of the meter is often required. A wide range of requirements for current meters is imposed on the functionality and method of recording the output signal: self-recording, digital, logarithmic, differential, etc. The operating conditions of low current meters can be varied: from laboratory to space.

- Charge sources

Objects that are adequate in their characteristics to the source of charges are electrophysical . In electrochemistry, electrobiology , etc., such objects do not occur.

The most common sources of electrical charges are piezoelectric transducers. They are used to measure pressures, forces and accelerations and use the direct piezoelectric effect. When determining accelerations on a piezoelectric transducer, a known additional mass is strengthened and the force, and thus the acceleration, is found from the resulting charge. Most often, the transducers are made of quartz, the piezoelectric constant of which is 2∙10 -12 Cl/N. Quartz transducers are not destroyed at pressures up to 4∙10 9 Pa and temperatures up to +500 o C , have a low temperature coefficient (0.01% / K ) and high insulation resistance (about 1 0 14 Ohm). They are linear over a range of about 120 dB with no hysteresis.

Quartz transducers are widely used in testing rockets and other large mechanical systems. At the same time, rather stringent requirements are imposed on charge meters. Input elements of meters can be located in close proximity to the objects under test or even together with transducers. However, it usually requires high

sensitivity, wide operating range and very wide frequency range. The latter is necessary in order to use the same meters to control both fast and very slow processes. The measurement error should lie at the level of 1 - 2%, and in some cases at the level of tenths of a percent. A number of specific requirements are imposed on the functionality of charge meters of piezoelectric transducers.

Another common type of object with an output signal in the form of a charge is the Faraday cell (cylinder, cage). It is a closed conductive screen isolated from surrounding objects. When charged particles or bodies enter the cell, their charge is transferred to the cell and can be measured. Faraday cells are used to measure the charges of individual particles, particle flows, flows of non-conductive liquids, in meteorological studies of precipitation, as well as in physical studies, in particular, in determining the electrification of liquid, granular and other non-conductive bodies. Since the range of problems solved with the help of a Faraday cell is very diverse, the requirements for the electrometric charge meters used in this case can vary greatly. In physical research, high sensitivity and accuracy are required; in meteorological work, one has to deal with field operating conditions.

Taking this opportunity, we note that in practical systems with a Faraday cell, not the charge of the cell, but its potential is often measured, and the charge is calculated from the values of the potential and capacitance of the cell. The attraction to such a more complex and less accurate method of determining the charge is explained by the fact that electrometric charge meters are not yet as widespread as electrometric voltage meters.

Charge meters are also widely used in various electrophysical experiments, when it is necessary to determine the integral value of the current over a long period of time, for example, o an ionization chamber, an ionization-flame detector of a gas chromatograph, a photocell, etc. This is required if it is not the power that is of interest, but the energy of the phenomenon under study, for example, the dose of X-ray radiation, light energy , etc.

It follows from the above consideration that objects adequate to charge sources are not as numerous as voltage sources or current sources. Nevertheless , depending on their purpose, almost the same diverse requirements are already being imposed on charge meters as on electrometric voltage and current meters. Often, the maximum possible sensitivity is required, and there is a need for meters of very high accuracy. When working with piezo transducers , it is important to have a wide operating and frequency range; operating conditions may be very different.

- Resistive objects

Measuring objects, which are characterized by high active resistance, are extremely widespread. These include all dielectrics and high-resistance semiconductors

in particular glassy, amorphous, organic and photoconductors . It is necessary to determine their resistance in physical experiments, as well as in the process of development, research and production of such materials. Typical is the problem of determining the volume and surface resistances of solid, liquid, pasty, etc., high-resistance materials.

Another typical task is to determine the quality of the insulation of products and equipment, such as cables, insulators, electronic lamps, semiconductor elements, electrical products. In physical experiments, naturally, one has to deal with more specific problems. Large resistance meters are widely used to determine the properties of the atmosphere in meteorological work. Recently, the problem of measuring the resistance of high resistance resistors used in electrometric equipment has become topical.

High resistance meters are the second most common electrometric instruments after current meters. The difference in their areas of use lies in the fact that high resistance meters, unlike low current meters, are not used to measure non-electric quantities, since modern technology does not have converters of non-electric quantities into high electrical resistance.

The requirements for high resistance meters are very diverse. When working with high-quality dielectrics, it is necessary to measure resistances up to 10 17 -10 18 ohms. The voltage across the object is usually high, from tens of volts to several kilovolts . However, when measuring the conductivity of some semiconductor objects, thin dielectric films, etc., a high voltage on the object is unacceptable. Sometimes a wide measurement range without switching limits is desired, then logarithmic meters are also used. In some cases, there are special requirements for meters, in particular, in devices for determining the leakage resistance of capacitors, special measures must be taken to exclude the influence of a large capacitance and an object on the operation of the meter. Usually, when measuring high resistances, an error of about 2 - 10% is allowed (depending on the measurement range). In some physical studies, even greater accuracy is required. In the control and especially in the study of high-resistance resistors, the error should not exceed tenths or even hundredths of a percent.

stringent requirements regarding the speed of response to electrometric resistance meters, since there are no high-resistance objects in which rapid changes in resistance can take place. Most often, the required speed is determined by the performance of the equipment, and not by the speed of the processes under study. The operating conditions of high resistance meters can be different: from laboratory to field.

- capacitive objects

Objects whose main characteristic parameter is capacitance are very common. In many cases, however, the capacitance of an object is independent of frequency and can be determined at sufficiently high frequencies, for example, in the audio range.

This category includes, in particular, all air containers. For measurements

capacitances in the audio frequency range, electrometric equipment does not

required. If the capacitance is represented by a capacitor with a dielectric, then, due to polarization and other phenomena, the dielectric constant, and thus the capacitance, depend on the frequency or time spent in this mode. Quite often, this dependence manifests itself precisely in the region of very low frequencies. The residual parameters of capacitive objects, which are characterized by active resistance or loss angle , may also depend on the frequency . Since such objects are, as a rule, high- resistive, electrometric equipment is inevitably used to determine their parameters.

Most often, objects for which it is necessary to know the dependence of capacitance and resistance on frequency in the infralow frequency region are samples of dielectrics. For example, it is known that the electrical characteristics of polymers in the infralow frequency range can be used to judge their molecular structure and mechanical properties. Therefore, studies of such characteristics are of great importance in the development of new and the study of known materials. Usually, to solve problems of this kind, the same samples and electrodes are used as in the measurement of the resistivity of dielectrics.

The need to measure the infra-low frequency properties may also occur in other cases, in particular, when characterizing capacitors with a dielectric intended for use in low-frequency circuits, when characterizing high-resistance resistors intended for use in high-speed electrometric equipment, etc.

Devices for measuring the characteristics of high-resistance capacitive objects at infra-low frequencies are the least common of all types of electrometric equipment; they are not yet mass-produced, there are only individual developments for research purposes. With the help of such devices, both the capacitance and resistance of the object, i.e., both components of its complex resistance, are simultaneously measured, therefore, the methods for constructing meters for capacitive objects have much in common with the methods for constructing meters for parameters of alternating current circuits. The difference is the requirement for high current sensitivity, which predetermines the use of electrometric equipment.

The upper limiting frequency for which such meters are made is 20-100 Hz ; at high frequencies, the necessary measurements are performed using serial meters for the parameters of AC circuits. The lower limiting frequency is 0.1 - 0.01 Hz, at lower frequencies the time of one measurement inevitably becomes very long. In practice, there is a need for measurements at lower frequencies, up to 10 -5 Hz.

The measurement range of such instruments in terms of resistance is usually the same as for meters of high resistance at direct current, i.e., on the order of 10 6 -10 16 ohms. The capacitance range is 1–1000 pF. By choosing the dimensions of dielectric samples, in most cases it is possible to fit into this range for any values of the permittivity of the sample material encountered in practice. Since such devices are still used exclusively for. research purposes, high demands are placed on their accuracy. This is also related to the fact that

the expected changes in capacitance and resistance with frequency may not be very significant. Typically, the measurement error of capacitance and resistance does not exceed 0.5 - 2%. So far, the operating conditions of the meters under consideration are always laboratory ones.

1-3. INDIRECT INTERFERENCE OF ELECTROMETRIC OBJECTS AND MEASUREMENTS

Rice. 1-4. Equivalent Circuit for Noise Calculation

R C-chains.

a) General remarks

When measuring signal parameters or circuit parameters of electrometric objects, one has to deal with the object's own noise. Their level determines the requirements for the sensitivity threshold of the measuring equipment. Any electrometric device also has its own interference, the level of which depends on the properties of its elements (electrometric voltage meter, passive electrometric transducers and structural elements ) and determines its limiting capabilities. Both objects and input circuits of the electrometric

equipments have the same defining features (high resistance and low current level), and the noise sources in them are of the same nature. Let's take a look at these sources.

b) Thermal noise

All elements of circuits with active resistance are characterized by thermal noise. It is caused by disordered thermal motion of charge carriers (electrons in metals and n - type semiconductors). The resistances of electrometric circuits are high; therefore , the emf is also high. thermal noise. Examples of high-resistance circuits with a high emf level. thermal noise can be electrometric voltage sources, resistive objects and high-resistance resistors.

Consider the thermal noise of the simplest R C-circuit (Fig. 1-4). Mean square emf thermal noise resistor with resistance R in the frequency range dω

where K is the Boltzmann constant; T is the absolute temperature.

The maximum value of the thermal noise voltage (peak-to-peak or, as it is sometimes said, peak-to-peak) is assumed to be 6 to 8 times the rms value. Fluctuations with a larger amplitude are extremely rare and can be ignored.

Noise voltage across the resistor. R , shunted capacitance C:

Rice. 1-5. Spectral distribution of noise in R C-circuit.

Rice. 1-6. Thermal noise of a resistive element with a large capacitance.

always ω 0 ≈0 (constant and slowly changing signals), and therefore

it follows (Fig. 1-5) that 99% of the total noise energy lies in the frequency range at which the condition 0≤ ωRC ≤100 is satisfied . With a further increase in ωRC there is no increase in the intensity of fluctuations. Reduction of thermal noise by narrowing the bandwidth ω or decrease in resistance R is achieved only when ωRC <5. For ωRC ≥10 we have:

28

where ω and ω 0 are the upper and lower limits of the frequency range in which noise is determined. In electrometry

Out of addiction

Рис. 1-7. Зависимость тепловых шумов от сопротивления и полосы пропускания.

29

and the noise voltage does not depend on R (Fig. 1-6) , namely

Ratio R ≥10∕ ωC often observed in electrometric circuits. At f =10 Hz and С =10 pF it is performed for R ≥ l ,6∙10 10 Ohm.

For ωRC << l we have arctg ωRC ≈ ωRC and then where ∆ f is the bandwidth in which the noise is determined.

This dependence is shown in fig. 1-7. Graphs built for R up to 10 14 ohms, and this requires explanation. Equation (1-1), and consequently, (1-4) at very high resistances cease to be valid, since in this case the spectral noise density tends to infinity, which cannot be. As shown in [L. 1-22 ], this apparent paradox arises only as a result of forgetting the fact that equation (1-4) is applicable only to sufficiently good conductors, i.e., those inside which, at a given frequency, the displacement current can be neglected compared to With conduction current. If this condition is not satisfied, then along with the conductivity, the permittivity of the conductor must also be taken into account . When taking into account, in addition to the resistance, the internal capacitance of the conductor, the following expression is obtained for the spectral density of the emf. thermal noise:

It follows from this equation that the spectral density at ω = const and C = const for small R grows in proportion to R , and for large R falls in proportion to 1 ∕ R . The maximum spectral density is reached at R max =1∕ ωC . For example, at C \u003d 1 pF and ω \u003d \u003d 6.3 1 / s (1 Hz) we have R max — 1.5* 10 11 Ohm.

For R = const and C = const the spectral density of thermal noise of high-resistance elements with increasing frequency from ω = 0 first does not depend on frequency, and then begins to decrease. When ω max =1/ RC the spectral density is halved. For high-resistance elements , in particular for insulators, this cutoff frequency is very low. For example, at R \u003d 10 15 Ohm and C \u003d 1 pF we have ω ma κc \u003d 0.001 s -1 , i.e. f max \ u003d 1.5 ∙ 10 -4 Hz.

Thus, the spectrum of thermal noise of insulators of electrometric equipment is concentrated in such a low frequency range (from 0 to 10 -4 - 10 -5 Hz), which is inoperative for such equipment. Indeed, the measurement time is always finite and does not exceed 100-1000 s; therefore, thermal noise with periods much longer than this time cannot affect the measurement results. They can appear only at measurement times commensurate with days (24 4≈10 5 c), for example, when recording the drift of an electrometer with a dynamic capacitor. In this case, the integral value of noise is determined by the equation ( 1-3 ).

It follows from the foregoing that the thermal noise of insulators in conventional electrometric equipment should not be taken into account and it must be assumed that the insulators do not have thermal noise.

- shot noise

The current is a statistical result of the movement of numerous electrons, which, in addition to the external electric field, also depends on other, random causes. Therefore, there are short-term deviations of the current from the average value, the relative value of which is the greater, the smaller the number of unit charges passes per unit time through the circuit, i.e., the smaller the current.

30

The mean square of fluctuations created by the shot effect in the frequency band ∆ f [L. 1-2]

where q \ u003d 1.6 * 10 -19 Cl is the electron charge, and the current I is equal to the arithmetic sum of the electronic and ionic components . Dependencies √ i 2 dr ( I , ∆ f ) according to equation (1-5) are shown in fig. 1-8. The double amplitude of the shot noise is 5–8 times higher than the root-mean-square value.

Рис. 1-8. Зависимость тока дробовых шумов от уровня тока и полосы пропускания.

Shot noise is taken into account when analyzing the sources of intrinsic interference of electrovacuum systems: electric lamps, electronic and photomultipliers, etc. In resistive circuits characterized by thermal noise, shot noise should not be taken into account. This follows from the law of uniform distribution, which is one of the main ones in statistical physics and according to which the average energy of fluctuations associated with each variable that enters the energy only through an additive quadratic term in the state of thermal equilibrium is equal to KT ∣ 2, where K is the Boltzmann constant, T is the absolute temperature.

- Parasitic current of insulators

In addition to fundamentally irremovable thermal and shot noise, there are other sources of intrinsic noise in real electrometric circuits. Any, even the most high-quality insulator gradually transfers the charges accumulated on the surface and in the volume to the electrometric circuit insulated by it, creating a parasitic current. This current is not associated with the presence of electrical voltage on the insulator. The reasons causing the parasitic current of the insulator are little studied, we can only tentatively list them. One of them is the polarization of the dielectric by a previously applied electrical voltage (Fig. 1-9). The current through the dielectric reaches the steady value I0 only some time after the voltage is turned on ( t = 0), which can be several hours. When the voltage is turned off ( t \ u003d t 1 ) , a current flows through the dielectric about

Table 1-1

Parasitic current of insulating materials under various conditions

| Material | Parasitic current, *10 -17 BUT |  |  |
| --- | --- | --- | --- |
|  | Normal conditions | Relative humidity 98% at 40 ° C | Temperature change from 20 to 70°C |
| Escapon | 2-3 | 20-30 | 20-30 |
| Amber | 2-20 | 100-300 | 20-30 |
| Fluoroplast-4 | 20-30 | 20-30 | 20-30 |
| Ceramics, varnished FG-9 | 20-50 | 200-600 | 20-30 |

Рис. 1-9. Зависимость тока через диэлектрик от времени.

military sign, and the time it takes for it to fall to zero can also be significant. The causes of parasitic current can also be electrification by friction, the piezoelectric effect under the influence of mechanical loads, electrochemical phenomena on the surface, a slow change in the structure of the material, etc.

The level of parasitic current depends on the material of the insulator , its volume (the larger the volume, the greater the current), design, mounting method, cleanliness of the surface treatment and its contamination , ambient temperature and, especially, the rate of its change, relative air humidity, time holding the insulator under steady state conditions, etc. In the best case, it is possible to reduce the current of the insulator to ( l ÷2)10 -17 And, at worst, it can be 10 -15 - 10 -14 A, and for a short time even more. In table. Tables 1-1 show the average values of parasitic currents for samples of 1 cm 3 in volume from different materials under different operating conditions. In prestressed insulators operating at pressure (2÷5)××10 6 Pa, parasitic currents are less than in non-voltage ones.

| Материал | Число α-частиц в 1 ч с поверхности 100 см2 |
| --- | --- |
| Медь, обработанная механически | 9 |
| Медь, выдержанная на воздухе | 21 |
| Медь электролитическая | 13 |
| Сталь коммерческая | 3 |
| Сталь нержавеющая | 1 |
| Латунь коммерческая | 5 |
| Кварц | 2 |
| Олово коммерческое | 121 |
| Олово химически очищенное | 14 |
| Припой | 2800 |
| Алюминий коммерческий | 31 |

D) Radioactive background

In gaseous electrometric objects (for example, an ionization chamber) and in the inlet air volumes of electrometric meters, parasitic current can be caused by background radioactive radiation. In this case, the ion density is low, and fields of the order of several millivolts per centimeter are sufficient for the directed movement of ions, which are created, for example, by a contact potential difference between insulated metal parts, which can reach up to 100 mV.

Table 1-2

Contamination of materials with sources of α - radiation

The radioactive background is called [L. 1-16] radioactive impurities in the materials from which parts are made, radioactive gaseous impurities in the gas (air) that fills the nodes of objects or meters, cosmic rays, radioactive radiation of the Earth and surrounding objects.

Materials for parts of objects and meters may contain radioactive impurities that give α- , β- , γ - radiation. The radioactive contamination of materials depends on the geological history of the formation of the deposit, the method of obtaining material from raw materials, the method of processing, etc. It can fluctuate over a wide range even for the same material. In table. 1-2 shows the characteristics of the average α -contamination of some materials [L. 1-16].

Steel, brass and quartz have the least activity. The background from α - radiation is due to the elements 3-377 	33

uranium-radium, actinium and thorium families and rarely earth elements. The energy of α - particles of natural sources lies in the interval E α \u003d (3.5 ÷ 17) 10 -13 J (2.2 - 10.5 MeV). Since the average energy required to create one pair of ions in the air is 5.3 * 10 -18 J

(33 eV), then at E α =8∙10 -13 J air ionization by one α - particle gives a charge of 8∙10 -13 ∙ l ,6 × 10 -19 ∕5.3∙ 10 -18 = 2.4∙ 10 -14 C, i.e., the parasitic current averaged over 1 hour is 2.4 ∙ 10 -14 ∕3600 = 0.7∙ 10 -17 A. From 10 to 100 α - particles appear in a conventional electrometric object or the input block of an electrometric meter in hour.

Materials can also be contaminated with β -emitting impurities , in particular K 40 , Pb 27 , H 3 , C 14 . Glass has the highest β -contamination [L. 16], with sing 50/50 and aluminum, the smallest - polyethylene, polystyrene and copper. Radioactive impurities in materials also give γ - radiation. The radioactivity of air or gas is determined mainly by the presence of radon, radiocarbon and tritium. Air radioactivity is 32 α -particles per hour per 100 cm 3 , argon - l , 3 ∙ 10 -3 , for nitrogen it is zero.

The main ways to reduce parasitic currents caused by background radiation are shielding , selection of pure materials, and rational design (reduction of ionization volumes, etc.).

- Electrostatic interference

The high resistance of electrometric circuits makes them sensitive to electrostatic interference. Any electrometric circuit is a parallel connection of high resistance and some capacitance, most often the time constant of this circuit is large. If an electrostatic field acts on the circuit, then when the field changes, a charge arises on it . The voltage it creates on the capacitance of the circuit appears as a false signal. Changes in electrostatic fields can be caused by the movement of charged bodies or by changes in the potentials of nearby surfaces. The movement of uncharged bodies can also have an effect if it leads to changes in capacitance and thus to a redistribution of charge. To protect against external electrostatic interference, elements of electrical

metric circuits are placed in a metal grounded screen. Sometimes it is beneficial not to ground the shield, but to connect it to some low-resistance point of the circuit, which is equipotential to the shielded section.

Electrostatic interference can also be caused by electric fields inside the screen, for example, created by charged surfaces of insulators. When the field changes or when high-resistance conductive elements move in the field, for example, due to vibrations, interference occurs. This is also caused by movements of conductive elements, such as mounting conductors, relative to each other or relative to the housing. Increased sensitivity of electrometric equipment to shocks and vibrations is typical. To eliminate this, rigid fastening of elements and mounting conductors, rational internal shielding, etc. are necessary.

- Contact potential difference

In electrometric circuits, not the contact difference itself is more often manifested, but its drift in time. In electromechanical devices, to reduce the drift of the contact difference, they resort to technology that ensures the maximum identity and invariance in time of the working surfaces (Chapter 9). The presence of a constant contact difference leads to interference if elements of high-resistance circuits move in the electric field created by it, for example, due to vibrations or due to design necessity (moving relay contacts, etc.). As mentioned, the contact difference can contribute to the appearance of a parasitic current in the presence of a radioactive background.

- Other sources of self-interference

In many elements of electrometric circuits , so-called flicker noise occurs . Their spectral density is approximately inversely proportional to the frequency, and at infralow frequencies the level of flicker noise is significant. Flicker noise largely determines the level of self-interference of high-resistance composite resistors, electrometric lamps, field-effect transistors, etc.

In a number of electrometric objects, specific sources of intrinsic noise are manifested, related to the principle of operation of these objects. Current fluctuations in the ionization chamber are caused by the discrete nature of the ionizing radiation. The relative rms value of these fluctuations is:

where Q is the total charge of ions of the same sign, created by one particle; I is the average value of the ionization current; R and C - load resistance and capacitance [L. 1-21].

In vacuum current sources (photocells, photomultipliers, etc.), additional noise is produced by field emission. In secondary emission sources (electron and photomultipliers), noise can be caused by fluctuations in the secondary emission coefficient. Secondary electron emission from the walls also affects the current fluctuations of ionization chambers [L. 1-46].

- External interference

One of the types of external interference (electrostatic) has already been mentioned, others include pulsed magnetic interference, interference from ground circuits, from power sources, and the background of the power network. The issues of combating external interference are not specific to electrometric equipment and you can use the recommendations for conventional electronic devices [L. 1-9].

- Limiting capabilities of objects and meters

Fundamentally unremovable intrinsic noise of electrometric objects and meters are only thermal or shot noise. Depending on the type of object or meter , these noises manifest themselves in different ways and one or the other must be taken into account.

Electrometric voltage sources have a finite internal resistance R i and, therefore, are characterized by a certain level of thermal noise. Depending on the ratio between the frequency band, 36

element, the limiting noise is determined by the thermal noise current of this element according to equation (1-6) and the noise of the voltage meter.

From this, in particular, it follows that attempts to create a high-quality source of low currents using a low- ohmic resistor and a low-voltage source cannot be fruitful.

The noise limit of current meters depends on the method used. In capacitor meters, there are no unavoidable sources of noise other than voltage meter noise, and storage capacitor noise is zero. In meters with resistive

Rice. 1-10. Dependence of the thermal noise current on the source resistance and bandwidth.

k) Evaluation of self-interference of electrometric objects

In real objects, the level of intrinsic noise significantly exceeds the level of unavoidable noise: thermal noise for voltage sources and resistive objects, and shot noise for vacuum current sources. The quality of an object can be characterized by its own noise factor [L. 1-12], equal to the ratio of all intrinsic interference of the object to their unavoidable part . The larger it is, the less sensitive electrometric equipment is required.

l) Evaluation of the sensitivity threshold of electrometric equipment

The quality of measuring devices in terms of sensitivity is usually evaluated [L. 1-2]

37

in which measurements are made and the source time constant, this level is determined by equations (1-2), (1-3) or (1-4).

An ideal vacuum current source has only shot noise. If the internal resistance of the current source is due to electronic conductivity and has a finite value R i , then it should be characterized by thermal noise (Fig. 1-10)

noise figure equal to the ratio of the total noise level at the output of the measuring device to that part of these noises, the source of which is the measuring object. The closer the noise figure is to unity, the less the meter contributes to the overall noise level. When determining the noise figure, one should take into account all the interference of the signal source, and not just their unremovable component. The noise figure has a minimum value with optimal matching of the signal source and the meter [L. 1-12].

Determining the noise figure is useful when making specific measurements where the characteristics of the object are known. The limiting capabilities of the meter itself, regardless of the objects with which it can be used, are characterized by the sensitivity threshold, defining it as the level of intrinsic noise in the absence of a signal, reduced to the input of the meter. If the self-interference level of the meter depends on the output capacitance of the signal source, then different levels of interference are normalized for different capacitance values. The sensitivity threshold of a meter depends on its bandwidth, which must be taken into account when comparing different sensitivity meters . It is most correct to normalize the level of the variable component of the meter's own noise using the curve of their spectral density in a given bandwidth, as provided by GOST 8009-72 "Normalized metrological characteristics of measuring instruments". In practice, this approach has not yet been developed, since it is not provided with methods and means for determining the required characteristics.

Designers of electrometric instruments often evaluate them by comparing the real sensitivity threshold with the theoretical one (in the same bandwidth), i.e., the degree of approximation to the maximum possible characteristics. This allows you to compare devices of different types with different frequency ranges.